Field Safety Corrective Action Report
Philips Healthcare

[image: image1]
BU iXR
-1/4-
XCR603-140329
Report Form
Field Safety Corrective Action
Medical Devices Vigilance System (MEDDEV 2.12/1 rev 5)


V.05/07


	1. Administrative information Destination

	Name of national competent authority (NCA)

	See Unit affected list

	Address of national competent authority

	Date of this report

	Reference number assigned by the manufacturer

	FCO72200269   /  CAPA#: TW3509618

	Incidence reference number and name of the co-ordinating national competent authority (if applicable)

	Identify to what other national competent authorities this report was also sent

	See section 7


	2 Information on submitter of the report

	Status of submitter

	Manufacturer

	Authorised representative within EEA

	[image: image2.png]


      Others (identify the role):


	3 Manufacturer information

	Manufacturer name
	

	Philips Medical Systems BV
	

	Manufacturer’s contact person
	

	T. op het Veld
	

	Address
	

	Veenpluis 4-6,  P.O. Box 10.000
	

	Postal code
	City

	5680 DA
	Best

	Phone
	Fax

	+31 6 10420792
	

	E-mail
	Country

	Twan.op.het.veld@Philips.com
	The Netherlands


	4 Authorised representative information

	Name of the authorised representative

	The authorised representative’s contact person

	Address

	Postal code
	City

	Phone
	Fax



	E-mail
	Country


	5 National contact point information

	National contact point name

	Name of the contact person

	Address

	Postal code
	City

	Phone
	Fax

	E-mail
	Country


	6 Medical device information

	Class
	

	AIMD Active implants

MDD Class III [image: image3.png]


       MDD Class IIb MDD Class IIa

MDD Class I
	IVD Annex II List A IVD Annex II List B

IVD Devices for self-testing IVD General

	Nomenclature system (preferable GMDN)
	

	GMDN
	

	Nomenclature code
	

	37621
	

	Nomenclature text
	

	X-Ray Imaging System
	

	Commercial name/brand name/make
	

	Allura Xper Systems
	

	Model number
	

	722003, 722010, 722011, 722012, 722013, 722017, 722121, 722122, 722123, 722124, 722133 and 722134

	Serial number(s) and/or lot/batch number(s)
	

	See UAL
	

	Software version number (if applicable)
	

	N/A
	

	Manufacturing date/expiry date (if applicable)

	Accessories/associated device (if applicable)

	Notified body (NB) ID- number
	

	344
	


	7 Description of FSCA

	Background information and reason for the FSCA

	We have received a customer feedback where the Monitor Ceiling Suspension system fell to its lowest position.  In this specific occasion it collided with the table top.

The investigation initiated concluded that the cause of the failure was an assembly error of the Actuator of the MCS.



	Description and justification of the action (corrective/preventive)

· The investigation initiated concluded that the cause of the failure was an assembly error of the Actuator of the MCS. The MCS must be moved –manuel or motorically- before there is a chance that it drops to its lowest position and User, patient or bystander need to stand in a close position in relation to the MCS before a hazardous situation might occur.

· Philips will replace the actuator of all affected systems free of charge (Mandatory FCO72200269).

· It is expected that this corrective action will be available as of February 2015.

In the meantime, as to mitigate the risk, Philips will secure the MCS actuator of the affected systems with straps (containment action Mandatory FCO72200274).

· Philips representatives will contact the involved customers as soon as possible in order to schedule the implementation of this action.

	Advice on actions to be taken by the distributor and the user.

· Inform all possible system users.

· In order to avoid any risk for patients, users or bystanders we recommend the following until the containment action FCO72200274 or corrective action FCO72200269 has been implemented: When system in use, we recommend that the user does not:

· Position or move the Monitor Ceiling Suspension above the patient.

· Allow staff to stand under or close to the Monitor Ceiling Suspension.

· In addition, please avoid any unnecessary movement of the Monitor Ceiling Suspension and inform all possible System users.

	Attached please find

	[image: image4.png]


Field Safety Notice (FSN) in English FSN in national language

[image: image5.png]


Others (please specify):  Unit Affected List. (UAL)

	Time schedule for the implementation of the different actions

	These countries within the EEA and Switzerland are affected by this FSCA

	-
EEA and Switzerland:

[image: image6.png]


Austria
Belgium
Bulgaria

Cyprus

Czech Republic [image: image7.png]


Denmark
Estonia
Finland
[image: image8.png]


France
[image: image9.png]


Germany

Greece
Hungary
Iceland
Ireland
[image: image10.png]


Italy Latvia
Liechtenstein
Lithuania
Luxemburg

Malta

[image: image11.png]


Netherlands
Norway

Poland
Portugal

Romania Slovakia
Slovenia
[image: image12.png]


Spain
Sweden
[image: image13.png]


Switzerland

[image: image14.png]


United Kingdom

· Candidate Countries Croatia
Turkey

All EEA, Candidate Countries and Switzerland.

· Others:

	These countries outside the EEA and Switzerland are affected by this FSCA

	AU, BR, CA, CL, CN, IN, JP, MX, US, ZA.



[image: image15]
I affirm that the information given above is correct to the best of my knowledge. Signature:

T. op het Veld
Best

Name
City
Date

Submission of this report does not, in itself, represent a conclusion by the manufacturer and/or authorized representative or the national competent authority that the content of this report is complete or accurate, that the medical device(s) listed failed in any manner and/or that the medical device(s) caused or contributed to the alleged death or deterioration in the state of the health of any person.
8 Comments








